## Offset Errors
Offset Errors - Digital Voltmeter DVM1
Enter values
Calc results
SPECIFICATIONS AND RESULTS
|  | Specifications | Error Budget Results | Should be positive
| Range | Offset Error (V) | Offset Error | Margin % | P/F ?
| 4 | 0.015 | 0.0122 | 18.6667 | PASS
| 20 | 0.075 | 0.0618 | 17.6 | PASS
OFFSET ERRORS
Offset Errors |  | Error Source | Error Value |  | Analysis Node | Range | Offset Sensitivity | Offset Error | Gain from Vin to Va | Offset Error RTI |  | Worst Case Sum 4V | Worst Case Sum 20V
Circuit Block | Ref, Part | e | e | unit | Va | Vrange | S | ΔVoffset = e*S (V) | Ka = Va/Vin | ΔVoffset_rti = ΔVa / Ka |  | abs(ΔVoffset_rti) | abs(ΔVoffset_rti)
ADC | ADC | ADC Voff | 0.0098 | V | vadc1 | 4 | 1 | 0.0098 | 1 | 0.0098 | V | 0.0098
ADC | ADC | ADC Voff | 0.0098 | V | vadc1 | 20 | 1 | 0.0098 | 0.2 | 0.049 | V |  | 0.049
ADC | ADC | ADC Res | 0.0024 | V | vadc1 | 4 | 1 | 0.0024 | 1 | 0.0024 | V | 0.0024
ADC | ADC | ADC Res | 0.0024 | V | vadc1 | 20 | 1 | 0.0024 | 0.2 | 0.012 | V |  | 0.012
Atten | SW1 | SW1 Ioff | 1e-09 | A | v1 | 4 | 0 | 0 | 1 | 0 | V | 0
Atten | SW1 | SW1 Ioff | 1e-09 | A | v1 | 20 | 160000 | 0.00016 | 0.2 | 0.0008 | V |  | 0.0008
Atten | SW2 | SW2 Ioff | 1e-09 | A | v1 | 4 | 0 | 0 | 1 | 0 | V | 0
Atten | SW2 | SW2 Ioff | 1e-09 | A | v1 | 20 | 0 | 0 | 0.2 | 0 | V |  | 0
[12 empty] | WCA 4V | WCA 20V
[12 empty] | 0.0122 | 0.0618

## Gain Errors
Gain Errors - Digital Voltmeter DVM1
Enter values
Calc results
SPECIFICATIONS AND RESULTS
|  | Specifications | Error Budget Results | Should be positive
| Range | Gain Error (%) | Gain Error | Margin % | P/F ?
| 4 | 1.5 | 1.1998 | 20.0133 | PASS
| 20 | 1.5 | 2.7998 | -86.6533 | FAIL
GAIN ERRORS
Gain errors scale directly to all nodes
Gain Error |  | Error Source | Error Value |  | Analysis Node | Range | Gain Sensitivity | Gain Error | Worst Case Sum 4V | Worst Case Sum 20V
Block | Ref, Part | e | e | unit | Va | Vrange | S | ΔK/K = e*S (%) | abs(ΔK/K) | abs(ΔK/K)
Atten | R1 | R1 Tol | 1 | % | v1 | 4 | 0 | 0 | 0 |  | %
" | R1 | R1 Tol | 1 | % | v1 | 20 | 0.8 | 0.8 |  | 0.8 | %
Atten | R2 | R2 Tol | 1 | % | v1 | 4 | 0 | 0 | 0 |  | %
" | R2 | R2 Tol | 1 | % | v1 | 20 | -0.8 | -0.8 |  | 0.8 | %
ADC | U1, ADC | ADC Tol | 0.2 | % | ADCword | 4 | 1 | 0.2 | 0.2 |  | %
" | U1, ADC | ADC Tol | 0.2 | % | ADCword | 20 | 1 | 0.2 |  | 0.2 | %
ADC | U2, Vref | Vref_err | 0.05 | V | ADCword | 4 | -0.19996 | -0.9998 | 0.9998 |  | %
" | U2, Vref | Vref err | 0.05 | V | ADCword | 20 | -0.19996 | -0.9998 |  | 0.9998 | %
[9 empty] | Sum 4V | Sum 20V
[9 empty] | 1.1998 | 2.7998

## DVM1 Calc
DVM1
Gain, Sensitivity Calc Sheet |  |  |  | Enter values
Calc results
Circuit, Gains, Levels
Source R | RS | 10 | Max source R (assume battery)
Input Atten | R1 | 800000
| R2 | 200000
Switch | Ron SW1,2 | 2
Signal Gain | K (4V} | 1 | v1/vin = 1
| K (20V) | 0.2 | v1/vin = R2/(R1+R2)
Ref | Vref | 5 | Input Full-Scale
ADC bits | N | 10
ADC counts | Counts | 1024 | 2^N
OFFSET SENSITIVITY
S (How does each error source contribute to an Offset error)
U1 voff | S | 1 | S = vadc1 / voff
U1 resolution | S | 1 | S = vadc1 / vres
SW1 ioff 4V | S | 0 | SW1 On, no effect
SW1 ioff 20V | S | 160000 | S = vo / ibp = R1¦¦R2
SW2 ioff 4V | S | 0 | negliable effect
SW2 ioff 20V | S | 0 | SW2 On, no effect
GAIN SENSITIVITY
S (How does each error source contribute to a signal Gain error)
K = vo/vs Signal Gain
S =(ΔK/K)/(ΔR/R) Sensitivity of K to a change in R
S = 0.5 means that a 1% change in R causes a 0.5% change in the K.
Similarly, a 0.1% change in R causes a 0.05% change in the K.
R1 | K | 0.2 | R2/(R1+R2) | K is the circuit gain
20V only | K' | 0.19841 | K(R1*1.01) | K' is the gain with a 1% tol for S calc.
| S | -0.79365 | ((K'-K)/K) / 0.01 | S = (ΔK/K) / (ΔR/R) = ((K'-K)/K) / 0.01
R2 | K | 0.2 | R2/(R1+R2)
20V only | K' | 0.2016 | K(R2*1.01)
| S | 0.7984 | ((K'-K)/K) / 0.01
ADC | S | 1 | Gain error (Tol) directly scales output
Vref | S | 1 | Voltage error (Tol) directly scales output
Vref | K | 204.8 | Ncounts/Vref
| K' | 204.759 | Ncounts/(Vref+0.001)
| S | -0.19996 | ((K'-K)/K) / 0.001

## ADC Bits to V
DVM1
ADC Error Calculations |  |  |  | Enter values
Calc results
Ref | Vref | 5 | Input Full-Scale
ADC bits | N | 10
ADC counts | Counts | 1024 | 2^N
OFFSET ERRORS
Input Offset | voff (bits) | 2
| voff (V) | 0.0097656 | V = Bits/2^N*Vref
Resolution | vres (bits) | 0.5
| vres (V) | 0.0024414 | V = Bits/2^N*Vref
GAIN ERRORS
Gain Error | K (bits) | 2
| K (%) | 0.19531 | V = Bits/2^N*100%